# AUDIT DETAILS

| Auditee: | GATT Technologies B.V. |
| --- | --- |
| CBW ID: | 113361 |
| Address: | Transistorweg 5, 6534 AT Nijmegen, LI, NETHERLANDS |
| Project Handler: | Spychaj Alexander |
| Audit date: | 27-09-2025 to 27-09-2025 |
| Audit language(s): | EN-GB |
| Audit mode: | Offsite |

Proposed next regular audit date: Proposed next regular audit type:

AUDITED FACILITIES

CBW ID	Facility details

113361	GATT Technologies B.V.

Transistorweg 5

6534 AT Nijmegen

120130	GATT Technologies BV

Transistorweg 7

6534 AT Nijmegen

# AUDIT CRITERIA

European Regulations:

| Scheme | Audit Type | Certificate number |
| --- | --- | --- |
| Annex IX Chapters I and III - MDR (EU) 2017/745 | Initial stage 1 audit | G12 113361 0001 Rev. 00 |
| Quality Management System Requirements: Scheme | Audit Type | Certificate number |

EN ISO 13485:2016 (+AC:2018+A11:2021)	Initial stage 1 audit	Q5 113361 0002 Rev. 00

Defined processes and documentation of the auditee’s quality management system

| Auditor: | Nadine Helfert |
| --- | --- |
| Audit Team Participants: | Spychaj Alexander (Co-Auditor) |

AUDIT TIMES

Date	Audit from	Audit to

27-09-2025	8:30 a.m.	5:30 p.m.

Distribution list:

Audit Team members

Certification Body / Notified Body

Applicant/Company

# CHECKLIST

| Item: | Result |
| --- | --- |
| Checklist: Stage 1 |  |
| Opening meeting Mandatory opening meeting topics: |  |
| Introduce team: introduce the TÜV SÜD team (e.g., lead auditor, auditor, experts, observers), including a description of their roles. | Completed |
| Representative: explain that the audit team represents the applicable certification/notified body and is responsible for executing the audit and creating the audit report. | Completed |
| Sampling: explain that the audit will be conducted using methods generally based on sampling (see ISO 19011 Appendix A.6) but the stage 1 audit is primarily focused on review of readiness for stage 2. | Completed |
| Audit type: confirm that audit type and objectives are correct. | Completed |
| Audit criteria: confirm that the audit criteria (i.e., services / schemes of the audit) are correct. | Completed |
| Language: confirm the audit language(s). | Completed |
| Certificate scope: review and confirm scope of certification by outlining all schemes that are applicable to the audit. Also, read the proposed scope statement(s) if applicable from the Appendix A/B/C and confirm them with the client. | Completed |
| Client data: confirm the employee count, shifts, business units, and location with the data in the applicable Appendix A/B/C. | Completed |
| Point of contact: confirm client's point of contact, roles and/or the identity of any guides and observers used for the duration of the audit. | Completed |
| Access: confirm access to the resources and facilities required to conduct an effective audit. | Completed |
| Safety procedures: confirm any relevant work safety, security, or emergency procedures with the client. | Completed |
| Areas of concern: confirm the client's understanding of areas of concern and that these could lead to nonconformities at the next audit. | Completed |
| Termination: confirm the client’s understanding of the condition under which the audit can be terminated prematurely. This is the case if available audit evidence indicates that the audit objectives are unattainable or suggests the presence of an immediate and significant risk (e.g., safety-related). | Completed |
| Updates: explain that the client is kept informed of audit progress and any concerns throughout the audit. | Completed |
| Communication channels: confirm the formal communication methods/channels to be used between the audit team and the client (interviews, face to face, ICT, etc.). If applicable, add relevant communication channels/ICT methods in the comment field. | Completed |
| Closing meeting(s): confirm the date and time for interim end-of-day meetings and the closing meeting(s). | Completed |

| Item: | Result |
| --- | --- |
| Questions: provide the client an opportunity to ask questions. | Completed |
| Opening meeting notes: use the comment field to provide opening meeting notes if necessary. | N/A |
| Audit mode |  |
| The stage 1 audit was conducted fully on-site. | No |
| Justification for conducting partially or fully offsite: The onsite presence of the audit team was not needed because offsite review of QMS documentation was sufficient (e.g., audit has no impact to location or processes). | Yes |
| Justification for conducting partially or fully offsite: The client as well as location(s) were already well known by the stage 1 performing lead auditor based on previously performed onsite audits. | N/A |
| Justification for conducting partially or fully offsite: Other. Please provide details in the comment field. | N/A |
| QMS documentation review |  |
| Does the client change from MDD/AIMDD to MDR? Comment: Product was not certified under MDD | No |
| Is the list of products to be CE marked under MDR correct according to the information listed in Appendix A.1? Comment: Appendix A/B/C (GATT Patch; Product description DHF-01-QR-001) | Yes |
| Quality manual including: - QMS scope clearly indicating applicable standards, relevant statutory and regulatory requirements, product related activities, geographical overview of QMS activities in case of multi-site organizations - details of and justification for any exclusion or non-application - reference to the documented procedures of the QMS - description of the interaction and interfaces between processes of the QMS Please enter document id, title and revision/date in the comment field. Comment: D: SOP-QA-422 GATT Quality Manual Version 5 2022-08-3 Area of Concern: TNR-000681 | Fail |
| Role of the organization: Please enter the role of the organization in the comment field (e.g., manufacturer, authorized representative, sponsor, importer, distributor, service provider). Please enter document id, title and revision/date in the comment field. Comment: D: SOP-QA-422 GATT Quality Manual Version 5 2022-08-3: manufacturer | Completed |
| Quality policy: Please enter document id, title and revision/date in the comment field. Comment: D: SOP-QA-422 GATT Quality Manual Version 5 2022-08-3 | Pass |

| Item: | Result |
| --- | --- |
| Quality objectives and related targets: Please enter document id, title and revision/date in the comment field. Comment: D: SOP-QA-422 GATT Quality Manual Version 5 2022-08-3 | Pass |
| Organization charts including responsibilities and authorities: Please enter document id, title and revision/date in the comment field. Comment: D: SOP-QA-422 Quality Manual, Annex 1 | Pass |
| Management review: - procedure - last management review including minutes and action items Please enter document id, title and revision/date in the comment field. Comment: R: F-MN-002_Minutes_20220309 D: SOP-QA-003 Management Review Version 3 2022-03-30 Area of Concern: TNR-000682 | Fail |
| Document and records control: - procedure - retention periods for documents and quality records Please enter document id, title and revision/date in the comment field. Comment: D: SOP-QA-002 Controlled Documentation Version 2 2021-05-20 | Pass |
| Internal audit: - procedure - internal audit program demonstrating coverage of full QMS for all audited schemes - internal audit report and related corrective actions (at least one completed examples) Please enter document id, title and revision/date in the comment field. Comment: D: SOP-QA-008 Version 6 2022-08-25 D: QR-QA-001 GATT Audit Plan Version 8 2022-09-23 D: QR_QA-001 GATT Audit Plan R: IA-01-2022 Audit Report & CAP_QMS Documentation R: IA-01-2021 Production process R: IA-02-2021 Human Resources Showing correction action and according CAPA R: IA-02-2022 Warehouse Control • R: IA-03-2022 Lab Audit R: IA-04-2022 Internal Audit Report & CAP – Manufacturing Clearly states number of NCs and objectives are met R: IA-05-2022 Risk Management R: IA-06-2022 Management Area of Concern: TNR-000683 | Fail |

| Item: | Result |
| --- | --- |
| Factory layout (if applicable, per site): - production area per production technology - warehouse(s) - material flow Please enter document id, title and revision/date in the comment field. Comment: SOP-PS-004 Work environment, Version 3 (January 17, 2022) No separate Site Master File or qualitycontrolled document with a facility plan exist. The purpose of SOP PS-004 is primarily to describe the process to maintain the work environment needed to achieve conformity to product requirements. In section 7 areas are defined. In the appendix layout requirements are included Area of Concern: TNR-000684 | Fail |
| Operational details including: - areas in shift operation - employees in shift operation (outside regular office hours) - subcontracted / outsourced processes Please enter document id, title and revision/date in the comment field. Comment: Opening Presentation Organigram / D: SOP-QA-422 Quality Manual, Annex 1 As of September 2022, the Nijmegen facility is at 16 employees, Naarden office 3 employees = 19 in total. There is a plan to hire 4-6 new employees in 4QUpdate of Company Profile | Pass |
| Validation of the application of computer software used in the quality management system: Please enter document id, title and revision/date in the comment field. Comment: D: SOP-QA-422 GATT Quality Manual Version 5 2022-08-3 | Pass |
| Human resources: Please enter document id, title and revision/date in the comment field. Comment: R: IA-02-2021 Human Resources Showing correction action and according CAPA | Pass |
| Infrastructure: Please enter document id, title and revision/date in the comment field. Comment: D: SOP-QA-123 Infrastructure V.1.0 | Pass |

| Item: | Result |
| --- | --- |
| Work environment: Please enter document id, title and revision/date in the comment field. Comment: # SOP-PS-004 Work environment, Version 3 (January 17, 2022) No separate Site Master File or qualitycontrolled document with a facility plan exist. The purpose of SOP PS-004 is primarily to describe the process to maintain the work environment needed to achieve conformity to product requirements. In section 7 areas are defined. In the appendix layout requirements are included TNR-000684 Fo AoC see | Fail |
| Contamination control: Please enter document id, title and revision/date in the comment field. Comment: D: SOP-QA-124 Pest Control V.1.0 | Pass |
| Risk Management: Please enter document id, title and revision/date in the comment field. Comment: R: F-QA-002 Risk Management Table Template referenced D: SOP-QA-004 Risk Management Procedure, V1 (September 24, 2020) Procedure in place. Commitment to apply risk management throughout the life cycle. Top management is responsible for defining and documentation of a policy for establishing criteria for risk acceptability. There are various techniques that can be used for Risk Analysis (PHA, FMEA, FMECA, HAZOP, HACCP, and others). The recommended standard technique for using in GATT risk analyses is FMEA. As per defined requirements, known and foreseeable risks are defined. Reasonably foreseeable misuse considered. Area of Concern: TNR-000685 | Fail |
| Customer-related processes including product requirements and communication with customers: Please enter document id, title and revision/date in the comment field. Comment: D: SOP-PS-027 Customer Order Process see TNR-000681 | Fail |
| Design and development including transfer of design and development and control of design and development changes: Please enter document id, title and revision/date in the comment field. Comment: D: SOP-PS-001 Design Control see TNR-000681 | Fail |
| Purchasing: Please enter document id, title and revision/date in the comment field. Comment: SOP-SC-001 Purchasing Version 3 20209-22 Area of Concern: TNR-000686 | Fail |

| Item: | Result |
| --- | --- |
| Verification of purchased product: Please enter document id, title and revision/date in the comment field. Comment: SOP-SC-001 Purchasing Version 3 20209-22 | Pass |
| Production and service provision: Please enter document id, title and revision/date in the comment field. Comment: D: SOP-PS-001 Design Control D: SOP-PS-027 Customer Order Process | Pass |
| Installation and servicing: Please enter document id, title and revision/date in the comment field. Comment: D: SOP-PS-001 Design Control D: SOP-PS-027 Customer Order Process | Pass |
| Validation of processes for production and service provision including validation of the application of computer software: Please enter document id, title and revision/date in the comment field. Comment: D: SOP-PS-001 Design Control D: SOP-PS-027 Customer Order Process | Pass |
| Validation of sterilization and sterile barrier systems: Please enter document id, title and revision/date in the comment field. Comment: D: SOP-PS-001 Design Control D: SOP-PS-027 Customer Order Process | Pass |
| Product identification: Please enter document id, title and revision/date in the comment field. Comment: D: SOP-QA-020 UDI and Interaction with Eudamed, Ver 0 (August 30, 2021 Area of Concern: TNR-000687 | Fail |
| List of devices in scope and status of marketing clearance: Please enter document id, title and revision/date in the comment field. Comment: D: SOP-QA-016 Regulatory Strategy and Medical Device File | Pass |
| Traceability: Please enter document id, title and revision/date in the comment field. Comment: D: SOP-6585 Travebility v1.0 | Pass |
| Customer property: Please enter document id, title and revision/date in the comment field. Comment: D: SOP-PS-027 Customer Order Process | Pass |

| Item: | Result |
| --- | --- |
| Preserving the conformity of product: Please enter document id, title and revision/date in the comment field. Comment: D: SOP-85236 Preservation of product | Pass |
| Control of monitoring and measuring equipment including validation of the application of computer software: Please enter document id, title and revision/date in the comment field. Comment: D: SOP-123547 Measuring equipment v.1.0 | Pass |
| Feedback: Please enter document id, title and revision/date in the comment field. Comment: D: SOP-QA-013 Vigilance system Version 2 2022-06-22 D: SOP-QA-012 Post Market Surveillance Version 2 2021-10-15 7. | Pass |
| Complaint handling: Please enter document id, title and revision/date in the comment field. Comment: D: SOP-QA-013 Vigilance system Version 2 2022-06-22 D: SOP-QA-012 Post Market Surveillance Version 2 2021-10-15 7. | Pass |
| Reporting to regulatory authorities: Please enter document id, title and revision/date in the comment field. Comment: D: SOP-QA-013 Vigilance system Version 2 2022-06-22 | Pass |
| Monitoring and measurement of product: Please enter document id, title and revision/date in the comment field. Comment: D: SOP-QA-422 GATT Quality Manual Version 5 2022-08-3 | Pass |
| Control of nonconforming products: Please enter document id, title and revision/date in the comment field. Comment: D: SOP-QA-006 Non-conformities, Corrective and Preventive Actions | Pass |
| Analysis of data: Please enter document id, title and revision/date in the comment field. Comment: SOP-QA-015 Measurement Analysis and Improvement – CAPA | Pass |
| Corrective and preventive actions: Please enter document id, title and revision/date in the comment field. Comment: SOP-QA-015 Measurement Analysis and Improvement – CAPA | Pass |

| Item: | Result |
| --- | --- |
| Other QMS documents relevant for stage 1: Please enter document id, title and revision/date in the comment field. Comment: D: SOP-QA-002 Controlled Documentation Version 2 2021-05-20 D: SOP-QA-008 Audits | Pass |
| MDR Article 10 (9) requirements – The quality management system has been established according to the applicable MDR conformity assessment procedure(s) and does address at least the following aspects. Evidence for implementation was provided in the listed documents: | Completed |
| a) a strategy for regulatory compliance Please enter document id, title and revision/date in the comment field. Comment: D: SOP-QA-016 Regulatory Strategy and Medical Device File | Pass |
| b) identification of applicable general safety and performance requirements (GSPR) Please enter document id, title and revision/date in the comment field. Comment: D: SOP SOP-QA016 “Regulatory Strategy and Medical Device File”, Vers 2 (July 13, 2022) R: F-QA-025 General Safety Performance Requirements Checklist, Vers 0 (August 30, 2021) Area of Concern: TNR-000688 | Fail |
| (c) responsibility of the management Please enter document id, title and revision/date in the comment field. Comment: D: SOP-QA-422 Quality Manual Version 5 chapter 5 | Pass |
| (d) resource management, including selection and control of suppliers and sub-contractors Please enter document id, title and revision/date in the comment field. Comment: D: SOP-SC-002 V3 Supplier Selection and Assessment, | Pass |
| (e) risk management as set out in in MDR Section 3 of Annex I Please enter document id, title and revision/date in the comment field. Comment: D: SOP-QA-004 Risk Management Procedure, V1 (September 24, 2020) See TRN-000685 | Fail |
| (f) clinical evaluation in accordance with Article 61 and Annex XIV, including PMCF Please enter document id, title and revision/date in the comment field. Comment: D: SOP-QA-014 Clinical Evaluation and PMCF, Vers 0 (July 14, 2021) Area of Concern: TNR-000689 | Fail |

| Item: | Result |
| --- | --- |
| (g) product realisation, including planning, design, development, production and service provision Please enter document id, title and revision/date in the comment field. Comment: D: SOP-PS-001 Design Control D: SOP-PS-027 Customer Order Process See see TNR-000681 | Fail |
| (h) verification of the UDI assignments made in accordance with Article 27(3) to all relevant devices and ensuring consistency and validity of information provided in accordance with Article 29 Please enter document id, title and revision/date in the comment field. Comment: D: SOP-QA-020 UDI and Interaction with Eudamed, Ver 0 (August 30, 2021) See see TNR-000681 | Fail |
| (i) setting-up, implementation and maintenance of a post-market surveillance system, in accordance with Article 83 Please enter document id, title and revision/date in the comment field. Comment: D: SOP-QA-012 Post Market Surveillance Version 2 202110-15 Area of Concern: TNR-000690 | Fail |
| (j) handling communication with competent authorities, notified bodies, other economic operators, customers and/or other stakeholders Please enter document id, title and revision/date in the comment field. Comment: D: SOP-QA-013 Vigilance system, Ver 2 (June 22, 2022) Area of Concern: TNR-000691 | Fail |
| (k) processes for reporting of serious incidents and field safety corrective actions in the context of vigilance Please enter document id, title and revision/date in the comment field. Comment: D: SOP-QA-013 Vigilance system, Ver 2 (June 22, 2022) Area of Concern: TNR-000692 | Fail |
| (l) management of corrective and preventive actions and verification of their effectiveness Please enter document id, title and revision/date in the comment field. Comment: D: SOP-QA-015 Measurement Analysis and Improvement – CAPA D: SOP-QA-006 Non-conformities, Corrective and Preventive Actions | Pass |
| (m) processes for monitoring and measurement of output, data analysis and product improvement Please enter document id, title and revision/date in the comment field. Comment: D: SOP-QA-025 Data Trending Procedure | Pass |

| Item: | Result |
| --- | --- |
| The client trained the relevant staff in terms of MDR. Please enter document id, title and revision/date in the comment field. Comment: F-RM-002 Annual Training Plan 2022 F-RM-003 Version 3 from 2021 MDR Certificate - Bram Keereweer.pdf MDR Certificate - Joost Opsteen.pdf MDR Certificate - Rosa Félix Lanao.pdf MDR Certificate - Stuart Head.pdf | Pass |
| The client conducted an internal audit covering all applicable requirements of the MDR. Please enter document id, title and revision/date in the comment field. Comment: SOP-QA-008 Audits QR-QA-001 GATT Audit Plan Internal audit reports: IA-01-2021 Production process IA-02-2021 Human Resources IA-01-2022 QMS and Document Control IA-02-2022 Warehouse Control IA-03-2022 Lab Audit IA-04-2022 Manufacturing IA-05-2022 Risk Management IA-06-2022 Management | Pass |
| The client conducted a management review focused on compliance to the MDR. Please enter document id, title and revision/date in the comment field. Comment: SOP-QA-003 Management Review F-MN-002_Minutes_220309 | Pass |
| Is the manufacturer established outside of a Member State and needs an authorised representative? | No |

| Item: | Result |
| --- | --- |
| Article 10: Technical documentation (TD) - TD according to Annex II and III - Retention period of 10 years for non-implantable devices- Retention period of 15 years for implantable devices Please enter document id, title and revision/date in the comment field. | Fail |
| Article 10: Technical documentation (TD) - TD according to Annex II and III - Retention period of 10 years for non-implantable devices- Retention period of 15 years for implantable devices Please enter document id, title and revision/date in the comment field. Comment: D: SOP-QA-016 Regulatory Strategy and Medical Device File, V02 (July 13, 2022) Technical File according to EUMDR, Annex II (Medical Device File according to ISO 13485) and a Summary Technical Dossier (STED) with attachments in order to submit to a Notified Body for CE approval and registration. Compiles by RA Manager. Based on a specific Template. SSCP mentioned (although not cross-referenced procedure). As per 6.10 TD is reviewed periodically (at least annually) as defined n quality planning. As per 6.11 archiving of TD and its revisions shall be archived according to SOP-QA-002. R: FA-QA-030 STED, Ver 1 (July 15, 2022) – Template with table of content defining in more detail the content of Technical DocumentationR: F-QA-025 Version 0 “General safety and performance requirements checklist (MDR)” August 30, 2021. A template from GSPR exists. The purpose of this document is to provide an overview of how general safety and performance requirements of the EU regulation 2017/745 are met for the product listed in the scope. D: SOP-QA-002 Controlled documentation Version 2 (May 20, 2021) GATT Technologies uses Greenlight Guru (GG) for document management. The retaining period for all GATT Quality Records is set for fifteen (15) years after the last device has been placed on the market. Area of Concern: TNR-000693, TNR-000694 | Fail |

| Item: | Result |
| --- | --- |
| MDR Article 10, 27, 29, 31: UDI and SRN - Procedure for UDI labeling including part B and C of Annex VI - Single registration number Please enter document id, title and revision/date in the comment field. Comment: D: SOP-QA-020 UDI and interaction with EUDAMED Version 0 (Aug 20, 2021) SRN registration is outlined. MDR Art 31: 6.2.2.1. An SRN has been registered in EUDAMED. Application forms issued by TÜV SÜD ensure that the SRN is given as part of the application. Checked with database entry and conformed. MDR Annex VI - UDI System: As per SOP-QA-020 section 6.2.3. Device Identifiers by GS1. Provides the basic UDI-DI using a format based on GS1 Global Model Number (GMN). Carrier: AIDC and HRI. The application of a UDI carrier on the device label is verified in F-QA016-1 Checklist label and IFU requirements). GATT will use specific label printer software, together with a generator software for designing and printing labels which, by design, provides a correctly formatted UDI carrier based on the input it receives. GATT will use specific algorithm and/or fonts for displaying a correctly formatted UDI carrier in a specific data matrix creator software MDR Art 29 - 6.2.4.1. Product Registration As per SOP-QA-020 section 6.2.4.1 of the procedure. MDR Article 27 (6) As per 6.2.3.1 GATT includes the Basic UDI-DI in the Technical Documentation (F-QA-030 STED template) and the Declaration of Conformity (F-QA-029). R: F-QA-016 Checklist Labeling and IFU requirements, Ver 2 (July 13, 2022) Template to document evidence how the labeling and instruction for use of specific product does meet the requirements; applicable for labels, implant cart, instructions how to use the implant cart, and the instruction for use. R: WI-QA-001 GATT-Patch Label Printing, Vers 1 (September 2, 2022) working instruction to provide a description of how to print labels and visually inspect them in such a manner that labels for GATT-Patch will be printed and checked in a reproducible manner. Includes instructions on UDI code labelling (generation and printing). Include only visibility check of barcode and UDI. Readability checked with a scanner on page 8. See TRN-000681 | Fail |
| Is the organization acting as an importer? | No |
| Is the organization acting as a distributor? Area of Concern: TNR-000695 | No |

| Item: | Result |
| --- | --- |
| Article 15: Person responsible for regulatory compliance (PRRC) (for guidance see MDCG 2019-7) - Appointment includes PRRC's responsibilities (15.3) - Sufficient qualification (15.1) - Manufacturer's Responsible Person (PRRC - 15.2) - PRRC needs to be an employee of the organization - Micro and small enterprises shall not be required to have the person responsible for regulatory compliance within their organization, but shall have such person permanently and continuously at their disposal - Organizations with more than one legal manufacturer under the parent company would need to ensure that each legal manufacturer has its own PRRC - for manufacturers located outside the EU, it must be assumed that the PRRC should also be located outside the EU - for manufacturers located in the EU, it must be assumed that the PRRC should also be located in the EU - Authorized Representative's Responsible Person (PRRC - 15.6) - Taking into account that the Authorized Representative is located in the EU, it must be assumed that any person to be permanently and continuously at its disposal should be also located in the EU - The PRRC for an authorized representative and for an 'outside EU' manufacturer cannot be the same person - The PRRC shall be responsible for ensuring (15.3) - The conformity of the devices is appropriately checked - The technical documentation and the EU declaration of conformity - The post-market surveillance obligations - The vigilance reporting obligations - For investigational devices, the statement is issued Please enter document id, title and revision/date in the comment field. Comment: R: JD – Rosa Felix Lanao (DOC1344) R: PRRC certificate, Medical Device RA certificate | Pass |
| Does the manufacturer handle systems and/or procedure packs? Comment: N/A found plausible. Within the scope of the assessment is an implantable device packaged in one unit packaging | No |
| Article 23: Parts and components - Items put on the market to replace are identical or similar to parts or components of a device Please enter document id, title and revision/date in the comment field. Comment: N/A found plausible. Within the scope of the assessment is an implantable device packaged in one unit packaging | Pass |
| Does the manufacturer have implantable and/or class III devices (other than custom-made or investigational devices)? Comment: N/A found plausible. Within the scope of the assessment is an implantable device packaged in one unit packaging | No |

| Item: | Result |
| --- | --- |
| Article 52, 20: CE marking - Procedure for CE marking Please enter document id, title and revision/date in the comment field. Comment: D: SOP-QA-016 Regulatory Strategy and Medical Device File, Version 2 (July 13, 2022) Section 6.7.1 includes Submission of TD to the NB CE not marking according to Annex V Area of Concern: TNR-000696 | Fail |
| Article 51, 52: Classification, Conformity Assessment Procedures - Classification rules applied? - Is the list of CE marked products up-to-date? - Are custom-made devices following procedure set out in Annex XIII (if applicable provide additional details in the comment field)? Please enter document id, title and revision/date in the comment field. Comment: D: DHF-01-SFT-183 Mode of action, qualification and classification GATT-Patch Provides only the documented evidence that the requirements have been fulfilled for the specific product within the scope of the assessment. N/A Custom-made devices are not within the scope of the assessment / certification (as per application). Area of Concern: TNR-000697 | Fail |
| Article 61: Clinical evaluation Clinical evaluation according to Part A of Annex XIV - Process in place / updated to generate and assess clinical data a) Example of clinical evaluation plan (CEP) - Identification of pertinent data - Appraisal of pertinent data - Analysis of clinical data - Clinical data report including PMS / PMCF plan - Example of PMCF plan Please enter document id, title and revision/date in the comment field. Comment: D: SOP-QA-014 Clinical Evaluation and PMCF, Version 00 (July 14, 2021) This procedure ensures that a clinical evaluation and post market clinical follow up (PMCF) are conducted in accordance with the European regulatory requirements and guidelines. Current MDCG guidance documents published / endorsed have been referenced and are implemented. R: F-PS-018-7 PMS Plan, Ver 2 (June 22, 2022: template to create a post-market surveillance plan Section 6.2.1 CEP references: F-PS-018-02 CEP Template F-PS-018-03 CER Template F-PS-018-04 PMCF Plan F-PS-018-05 PMCF report Section 6.6 PMCF The PMCF process starts with the establishment of a PMCF plan based on the outputs of the clinical evaluation. See the template. As soon as planned activities according to the PMCF-plan have been completed, the findings of the PMCF shall be analyzed and the results documented in a PMCF evaluation report. D : SOP-QA-012 Post Market Surveillance, Ver 2 (October 15, 2021) See TRN-000681 | Fail |

| Item: | Result |
| --- | --- |
| Article 62 to 82: Clinical investigation - Procedure for clinical investigations Please enter document id, title and revision/date in the comment field. Comment: D: SOP-QA-014 Clinical Evaluation and PMCF, Version 00 (July 14, 2021) Section 6.3.5 Clinical investigation Correctly defines that if no sufficient clinical data is available clinical investigations are to be considered. For the set-up of a clinical investigation, the requirements of local authorities in the region where the clinical investigation will be performed, have to be taken into account. The use of a Clinical Research Organization can be considered. ISO 14155 defines international requirements for performing a clinical study. Annex XV in MDR 2017/745 describes the requirements for clinical investigations in the EU. Area of Concern: TNR-000698 | Fail |
| Article 83 to 86: Post-market surveillance system - Procedure for post- market surveillance (Article 83) - Example plan for post-market surveillance (Article 84) - Class I device: Template for surveillance report (Article 85) - Class IIa, IIb and III device: Periodic safety update report (Article 86) - Update of technical documentation - Implantable and class III: Upload to Eudamed / notified body Please enter document id, title and revision/date in the comment field. Comment: D: SOP-QA-012 Post Market Surveillance, Version 2.0 October 15, 2021 Aspects included in the procedure, in particular in section 6.2. A standardized template exist (F-PS018-7). The initial PMS activities shall be planned before the first placing on the market of the product and updated as necessary during the product life cycle. PMS Plans is always updated once the after PSUR was submitted (i.e. annually). Art 85 - PMS report) PMS report described template available but within the scope of the certification there is only a class III product. Art 86 – PSUR Applicable based on the scope of the certification. Reporting requirements as per MDR Art 86 (2) defined in 6.3: both pathways via EUDAMED and directly included. D : SOP-QA-020 UDI and Interaction with Eudamed R: F-PS-018-7 PMS Plan, Ver 2 (June 22, 2022: template to create a post-market surveillance plan R: F-PS-018-6 - PSUR template _MDR, Version 0 (August 30, 2021) template to create a PSUR See TRN-00681 | Fail |

| Item: | Result |
| --- | --- |
| Article 87 - 89: Vigilance - Procedure - Serious public health risk: immediately but max. 2 days - Death: immediately but max. 10 days - Serious incident: immediately but max. 15 days - Trend reporting / periodic summary reports Please enter document id, title and revision/date in the comment field. Comment: D: SOP-QA-020 UDI and Interaction with Eudamed Provides link of EUDAMED with Vigilance reporting, Section 6.2.5: Are submitted to Eudamed manually through the Vigilance module. Vigilance reporting in Eudamed pertains to serious incidents, field safety notices and field safety corrective actions. Responsible: PRRC Vigilance. D: SOP-QA-013 Vigilance system Version 2 (June 22, 2022) The purpose of this Standard Operating Procedure is to provide guidelines for notification and evaluation of serious events or incidents and Field Safety Corrective Actions (FSCA) involving CE medical devices to the relevant authorities. Timely reporting of serious events or incidents and Field Safety Corrective Actions (FSCA) is a legal requirement. Responsible: PRRC Vigilance. 6.4.3: Incidents which occurred outside the EU do not require reporting to the national authorities, unless these incidents result in Field Safety Corrective Actions in a third country in relation to a device which is also legally made available on the Union market, if the reason for the field safety corrective action is not limited to the device made available in the third country. 6.8 Trend analysis and trend reporting D: SOP-QA-012 Post Market Surveillance, Version 2.0 October 15, 2021 MDR Art 88 Section 6.8 Trend reporting – referes then back to the vigilance procedure: In case complaints or incident trend results show a significant increase this might also lead to a report to the National Competent Authority (SOP-QA-013 Vigilance); Data Analysis: The company shall define frequency and time span of the PMS data collection, analysis and review. Identify and document the analysis tools and techniques that will be used to analyze the PMS data collected. Data analysis tools can be divided in two main groups based on the ability of using: quantitative (numerical) data; qualitative (descriptive) data; Verdict (pass, fail, n/a) Note: No specific / additional work instruction for FSCA / FSN or recall also detailing handling of returned products in place at current state. Seee TRN-00681 | Fail |
| Annex I 14.2 (d): IT security - IT security processes implemented (design, post-market operation) Please enter document id, title and revision/date in the comment field. Comment: N/A was found plausible. The manufacture is located within the EU. | N/A |
| Annex IX 2.2 - The manufacturer has defined quality objectives Please enter document id, title and revision/date in the comment field. Comment: D: SOP-QA-422 GATT Quality Manual Version 5 2022-08-3 | Pass |
| Summary of QMS documentation review The review of the QMS documentation has shown that the essential criteria are fulfilled. If the essential criteria are not fulfilled and no areas of concern are raised, please provide a justification in the comment field. |  |
| An updated application provided by the client is required to adjust the calculation based on the changes identified in the stage 1 audit. | No |

| Item: | Result |
| --- | --- |
| A revised offer provided by TÜV SÜD is required to adjust the calculation based on the changes identified in the stage 1 audit. | No |
| The audit team verified that the client has identified the relevant statutory and regulatory requirements and implemented these requirements in the QMS. | Yes |
| The review of the QMS documentation has shown that the essential criteria are fulfilled. If the essential criteria are not fulfilled and no areas of concern are raised, please provide a justification in the comment field. | Yes |
| Multi-site |  |
| Is this a single-site organization? | No |
| Is this a multi-site organization? | Yes |
| Are all sites operating under the same QMS? | Pass |
| Has the site/facility representing the central function responsible for the QMS been identified? | Pass |
| Are all sites/facilities covered by the client’s internal audit program? | Pass |
| Is the QMS subject to a centralized management review? | Pass |
| Has the central function the organizational authority to define, establish and maintain the QMS? | Pass |
| Is the central function ensuring that data is collected and analyzed from all sites? | Pass |
| Is the central function able to demonstrate its authority and ability to initiate organizational change as required in regard, but not limited to: - system documentation and system changes - management review - complaints - evaluation of corrective actions - internal audit planning and evaluation of the results - statutory and regulatory requirements pertaining to the applicable standard(s) | Pass |
| Add comments needed to understand the multi-site structure and site-specific activities if required in the comment field. | N/A |
| The criteria for a multi-site approach are fulfilled. If no, see areas of concern. | Yes |
| Are the sites from the multi-site organization sampled for auditing? | Yes |
| Are the sites performing similar processes/activities (e.g., sales, service) and are the sites performing processes/activities in the main technical area for "parts or services" in accordance with IAF MD9? | Pass |
| The criteria for a multi-site sample-based approach are fulfilled. If no, see areas of concern. | Yes |
| Audit results |  |

| Item: | Result |
| --- | --- |
| Based on the information of the application and the information collected during this stage 1 audit, it can be concluded that the allocation of resources for the onsite / stage 2 audit is ensured and that the client agreed to it. | Pass |
| Stage 1 audit conclusion and recommendation results are fully positive. | No |
| Stage 1 audit conclusion and recommendation results are partly positive. | Yes |
| Stage 1 for inital audit: Objectives of the stage 1 audit are essentially fulfilled, but areas of concern were identified that could be classified as a nonconformity during the stage 2 audit. The actions related to the areas of concerns shall be implemented prior to the stage 2 audit. | Yes |
| Stage 1 audit based on change notification: Objectives of stage 1 related the change notification are essentially fulfilled, but areas of concern were identified that could be classified as a nonconformity. The actions related to the areas of concern shall be implemented prior to the next audit. | N/A |
| Confirmation of next steps |  |
| Next audit date is planned and confirmed for stage 2 audit. Note: The proposed date is documented in the audit report. Comment: 03.11.2022 | Yes |
| Next audit date is planned and confirmed for recertification audit. Note: The proposed date is documented in the audit report. | N/A |
| Next audit date is planned and confirmed for surveillance audit. Note: The proposed date is documented in the audit report. | N/A |
| Next audit date is planned and confirmed for a special audit. Note: The proposed date is documented in the audit report. | N/A |
| Has the first round of all planned TD reviews been completed and deficiency reports been issued to the client? If no, ensure that the client is aware that first round of deficiency(ies) must be issued before stage 2 audit can be started. | Yes |
| Closing meeting |  |
| Thank you: thank the client for their hospitality and cooperation during the audit. | Completed |
| Uncertainty: explain that the objective evidence was obtained based on sampling, thereby introducing an element of uncertainty. | Completed |
| Areas of concern: present any identified areas of concern | Completed |
| Handling of areas of concern: explain the process for handling areas of concern, the timeframe of submitting the action plan and actions. | Completed |
| Item: | Result |
| Recommendation: explain that the recommendation for stage 2 is only given upon acceptance of the action plans for the areas of concern (if any). Also explain that the stage 2 audit is recommended in case there are no areas of concern. | Completed |
| Hand out: provide the client with the overview of areas of concern | Completed |
| Next audit: before leaving the client's facility, agree upon tentative dates for the next audit, regardless of the audit type Comment: 03.11.2022 stage 2 | Completed |
| Questions: provide the client the opportunity to ask questions, also after the audit ended | Completed |
| Complaint: inform the client that complaints can be submitted directly through you or the MHS complaint representative of your region/location. | Completed |
| Closing meeting notes: use the comment field to provide opening meeting notes if necessary. | Completed |
| Checklist: Stage 1 |  |
| Pre-audit TÜV SÜD INDEPENDENCE, IMPARTIALITY, CONFLICT OF INTEREST AND CONFIDENTIALITY (TO BE SIGNED PRIOR OPENING MEETING) |  |
| Auditor |  |
| I confirm for TÜV SÜD Product Service GmbH: Conflict of Interest & Consultant Activity Record (ID 2538) | Completed |
| Audit team participants |  |
| Confirmation by audit team participant - I confirm for TÜV SÜD Product Service GmbH: Conflict of Interest & Consultant Activity Record (ID 2538) | Completed |

# Disclaimer Statement

Auditing is based on a sampling process of the available information and therefore is not a guarantee of 100% conformity with requirements. Any audit recommendations are subject to an independent review prior to a decision concerning the awarding, continuation, update, or renewal of certification. A management system certification audit (initial, surveillance or recertification audit) is not a legal compliance audit (see ISO/IEC 17021-1:2015, 9.2.1.2).

Date, Auditor

This report may only be quoted in full. Any use for advertising purposes must be approved in writing. This report is confidential and the property of TÜV SÜD Product Service GmbH.

# Version History

| No. | Date | Name | Description of change |
| --- | --- | --- | --- |
| 1 | 2025-11-03 | Nadine Helfert | - Initial issuance / Completion of record |